This is the beginnings of a document describing basic cleanliness procedures for the 40m Lab, including maintaining cleanliness while opening chamber doors during vents.  It is referred to as “T10XXXX” in the Crane Safety Document.

Right now it’s kind of a mess, but I’ll clean it up and make it sensible ---Jenne

1, clean up surrounding area for good accessibility 2, wipe chamber and doors with methanol soaked "Kimwipes" 3, check 0.5 micron particle counts, it should be below 10,000 particles/cf min 4, turn on all HEPA filter units in this room to minimize this count,

  - except if it is blowing directly into the door that will be opened,

Door holder stands should be covered with clean CPStat100 (“AmeriStat”) before placing chamber door(s) in the stands.

Also discuss supplies needed for opening door: light aluminum door, kimwipes, lint free wipes, etc, etc,

7, vent annulos grove of the chamber involved,

  - this can be done on MEDM screen /cvs/cds/caltech/c0/ve/VacControl_Bak.adj please be wear that the vacuum screen is fully manual and require designated-authorized operator

8, minimize this vacuum control screen after venting in order to avoid accidental switching of vac valve 9, prepare clean CPSat100 and tape for temporary door cover, folded insides must be kept clean

10, bring 4 pieces of 3/4" open end wrench 11, bring 6 pieces of dry 1/2-13 x 2" hex head bolts, 12, bring lint-free wipers, 2 boxes of Kimwipes and Methanol in 2 spray bottles 13, bring O-ring cover ring and wipe it with methanol soaked "Kimwipes" 14, bring clean room gloves 15, bring light Al door and wipe it with Methanol soaked Kimwipes

2, close main laser shutter and place 2 manual blocks on PSL table, 3, close shutter and block other aux lasers like "green locking" and or "length measurement NPRO" ...or/and

  - other light

22, wipe exposed O-ring with dry lint free wipe and cover it with O-ring cover or light Al door

23, use 4-6 dry- clean 1/2-13 x2" bolts and nuts to secure cover and light door by finger torque

24, to maintain chamber cleanliness it is necessary to minimize chamber open condition 25, if you removed a door, you are responsible for closing it