40m Crane Safety Document: Corner-Vertex Area
 Steve Vass, Rana Adhikari, Jenne Driggers

Last Updated: 27 August 2010 Steve Vass

LIGO Document Number T10XXXX
Scope
This document describes the training procedures and general safety policies to be observed while utilizing the cranes and lifting equipment at the 40m Prototype Laboratory.  It also describes the procedures for removing and replacing the ~200lb steel vacuum chamber doors, which are the most common use of the cranes in the 40m Lab. This document refers specifically to the folding I-beam crane at the Corner-Vertex area of the Lab.  All 40m staff members who will work with or near the cranes must read and understand the entirety of this document. 

Description of the Crane
· 1 Ton capacity David Round / Hercules Wall Cantilever Jib Crane with Foldable I-beam
· This crane has the following features:

· 10 ft. I-beam with a foldable 4 ft. “swing” section

· When 14 ft. span is used the I-beam latch has to be locked

· 180 degrees of rotation of the main I-beam

· Main beam is at 11 ft. hub height

· 480V 3-phase power with manual disconnect that can be locked out

· Model A1-D1 LO-HED, Type A overhead wire rope electric hoist

· Updated with VDF, variable frequency driven control

· This modification was implemented 18 March 1992 by Rosa Automation Engineering, Inc.

· For most applications, we use a Lift All 12" nylon sling with 3200 lbs capacity

Requirements for Operation of the Cranes
The following are required before any crane may be energized or operated:

· A step-by-step procedure manual (this document)

· An updated crane certificate and periodic inspection reports

· Periodic inspections shall be conducted at least four times a year.

Three of them can be done in-house by Steve Vass or Bob Taylor or an other certified operator (according to ???, which OSHA reg ???)

· One of the 4 inspections has to be done by a certified crane inspector.

No load test is required.

                  -     Crane certification must happen immediately after repair, part replacement

or crane upgrade at full load by certified crane inspector. No annual certification is required.

· Unit must be locked-tagged out if not certified or inspected as listed

· 3 people present, one of whom is the designated Lead Crane Operator

· Minimum requirement is one 40m experienced crane operator with 2 cognizant assistants (i.e. authorized crane operators from other sites or trainees

Pre-Operation Checklist

This must be done in written form.  Print this checklist, check off items as completed.  Sign and date the checklist, and log it into the Crane Folder before loading the crane.

If the crane fails any inspections / tests, stop work immediately and notify Steve Vass.

Date: ___________ Name of Lead Crane Operator: ______________________________

Crane to be operated:                Corner-Vertex

1. ( Steel Toe Shoes are worn by Lead Crane Operator and all personnel in the vicinity of the crane

2. ( Laser Safety Goggles worn by all personnel.  Plastic goggles only, no KG5 glass safety goggles

3. ( Visual Inspection of crane and area

a.    Adequate room lighting

b. Inspection of overhead area – ensure that nothing is in a position to fall on personnel or equipment

c.    Inspection of area below crane and load path – ensure that nothing is blocking any walking paths, and that load path is clear, including rotation of crane

d. Initial inspection of crane

i. No dripping oil

ii. Sling is not loose

iii. No exposed wires

iv. Cable has fewer than 10% of strands damaged

v. Crane is clean.  If it seems dirty, wipe it down.  This must be done with at least 2 people since it requires climbing a ladder to reach the top of the crane.

4. ( Confirm that trolley is on the main I-beam within 2-4ft of vertical I-beam

5. ( Energize the crane by turning on the power at the wall switch
6. ( Turn the crane on by pushing the green Start button on the yellow hand-held control unit and raise load-block to 7-8 ft height.

7. ( Move to a safe distance from the crane, ~4-6 feet from the main beam. Do not ever stand underneath the crane

8. ( Test drive the trolley toward inside stop, observe the limit (~2ft from rotational axis)
9.  ( Test folding and locking of I-beam from folded position of the swing section

a. Rotate I-beam to position where there is sufficient space: at least 15ft available in the horizontal plane

b. Hold main I-beam by yellow umbilical while another person pulls hard on rope labeled “swing section” to straighten the 4ft extension

c. Listen for loud mechanical “click” of the latch as it engages all the way down

d. Confirm by eye that the latch is locked in and  it is fully down

10. (  Drive the trolley out slow/fast/slow to ~14ft until it reaches the stop. This is   maximum reach. Ensure you are in an area of clear of equipment.

11. ( Test drive jib down to shoulder height to inspect load-block (2 pulleys, swivel hook and hook latch)
a. Check 2 lock nuts on each side of load-block

a. Check for wear on the pulley wheels

b. Drive cables through and observe cleanliness and well-balanced horizontal position

c. Check swivel hook for cracks and deformation

d. Check latch spring for firm locking

12. ( Test drive load-block slowly and fast, feel half-depressed vs. fully-depressed speed

                        a.   Check steel wire rope for twisting, kinking, wear or broken strands

b.  Check that the wire rope seats in the pulley sheaves smoothly

c. Wipe steel wire rope with lint free wiper to check for residue

13. ( Test drive load-block upward to check for proper parallel reeving on the drum. Ensure that the wire rope never crosses over itself

14. ( Test drive load-block gently upward into limit switch chain.  Ensure that this disables upward drive, while allows downward drive

15. ( Test drive variable speed of load-block. Check that both slow (half-depress the button) and fast (fully depress the button) speeds work

16. ( Inspect sling for wear

a. Inspect sling for corner wear

b. Inspect sling for tears, snags, embedded particles, broken stitching, worn stitching

c. If there is any condition which causes doubt as to the strength of the sling, do not use it.  Replace the sling before applying any load.

17. ( Inspect the shackle and shackle bolt for cracks, nicks, deformation and wear

18. ( Place sling into hook, latch hook

19. ( Drive load block up to 7-8 ft height and drive trolley in onto main I-beam, ~2ft from vertical I-beam

20. (  Test folding and unlocking from straight locked I-beam position

a. Rotate I-beam to position where the 4ft swing section has free space to move

b. Pull down on latch rope labeled “folding latch open”. This will open the latch and the swing section of the I-beam can be folded 

c.  Hold main I-beam by yellow umbilical while another person pulls horizontally on rope labeled “swing section” until the angle between the swing section and the main beam is 70-80 degrees

Completing this checklist confirms that you have tested all operational functions of this 1 ton crane without load.  You acknowledge that the Lead Crane Operator has prime responsibility for the safety of all personnel and equipment.

Signature of Lead Crane Operator ________________________________ Date _______

The following sections describe the procedures for safely removing a chamber door from one of the test mass chambers.  As this is the most common application of cranes in the 40m Lab it is important for all crane operators to be comfortable with these procedures.

The load to be lifted with the crane is typically a side door of a chamber of the vacuum envelope.  These doors are 1 inch thick, 35 inches in diameter, and approximately 200lbs.  They are all attached (during normal operation) to the chamber with 16 bolts.  Each door has 2 critical areas which must be carefully watched during all lifting and moving operations: the finely ground surface near the outer diameter of the inside of the door which is used for sealing the O-rings, and the glass view ports.  Care must be taken to protect these super flat surfaces and the glass windows to ensure safe and functional operation of the 40m Interferometer.

While not directly in the scope of this document, it is important while working in the 40m Lab to maintain basic cleanliness procedures.  Please see T10XXXX for these details.
Pre-Loading Preparations
1. Position door holder stand within the crane’s reach (this is where the load will be placed)

2. If door will need to be moved to a different area, bring hydraulic lifter dolly (“Wesco”) and its door holder stand within the crane’s reach

3. Ensure all tools and supplies needed are easily accessible, including wrenches to loosen bolts securing door to chamber (see T10XXXX)

4. Ensure vacuum envelope and annuluses are at atmospheric pressure (“vented”). (See T10XXXX)

5. Re-think through entire plan (from loading the crane and lifting the door to securing the door in a stand), and confirm that it is in the 40m elog

a. What is the weight of this item? 

b. Where is the center of gravity of this item? 

c. Where is your attachment point? Let the crane hold it and gravity center it.

d. Recall you are planning and guiding this operation. 

6. Go through Pre-Operation Checklist, log it in the crane logbook
7.   Ensure you have at least the minimum number of personnel required: One authorized team leader (Lead Crane Operator) and 2 helpers

8.Ensure all personnel in the area have safety equipment on, including steel toe shoes

Loading the Foldable Crane at Vertex Area with a 200 lbs Door from BS Chamber

1. Set I-beam to folded position: see Pre-Operation Checklist 19 and 20 if needed

2. Rotate I-beam to be directly above the load

3. Drive trolley above load and lower load-block so sling is touching the top of the door

4. Attach shackle to the door by screwing in the 3/8 inch bolt until finger tight.  The ground (smoothed) side of the shackle should face the chamber

5. Remove 14 of the 16 bolts attaching the door to the chamber.  Leave the 2 bolts at the 10 o’clock and 2 o’clock positions in, with nuts loosened.  Note that the sling tension should still be fairly loose, and the door’s load should be sitting on the 2 alignment pins at the 9 o’clock and 3 o’clock positions

6. Drive the trolley away from the chamber by about 2-3 inches such that there is a small angle on the steel rope.  This will facilitate pulling the door away from the chamber

7. Put tension on the sling by bringing the load-block up slowly.  The sling should be tight enough that pushing on it with your finger results in a maximum of 1-2mm of deflection.  Note that you want the load to be shared between the crane and the pins, not yet entirely supported by the crane.  If the sling tension is too high, you risk lifting the entire chamber, however if it is too loose you risk having the full load on the alignment pins.  An experienced person should confirm the tension is acceptable before removing the final 2 bolts.

8. A “Door Handler”, who must have prior experience as Door Handler, holds the door against the flange and chamber (not holding any load, just preventing the door from moving)

9. A helper removes the final 2 bolts

10. Door Handler announces “Door is coming off” and forcefully pulls on the door to remove it from the pins.  Door Handler ensures that door does not swing back and touch alignment pins or any other part of the chamber (to protect the surfaces). Note that if the sling tension was correct the door should only move horizontally, and not up or down

11. Door Handler requests Crane Operator to move trolley farther away from the chamber

12. Door Handler and / or helper keeps hands on the door to guide it and protect it from swinging or hitting anything as it travels to the storage stand

13. Prepare to fold and lock I–beam if you have to go 14ft. 

a. See Pre-Op Checklist item 7, but note that this time the straightening of the swing extension will be done with a load on the main I-beam.  

b. Ensure that the trolley and load is fully on main I-beam, and not at all on swing section during all latching / unlatching procedures

c. Pay special attention to guide and stabilize the suspended door.  It requires 2 in-training operators to protect it from swinging
d. Confirm that I-beam is fully latched before moving trolley onto folding section

14. Rotate I-beam to stand location and drive trolley so that door is aligned directly above stand

15. Lower door gently onto storage stand, while protecting the door from scratches

16. Remove shackle bolt from door

17. Clamp door with C-clamp to the storage stand to prevent tipping

18. Bring load-block to the fully up position, and drive trolley in toward the wall

19. Turn crane power off using the hand-held control box first

20. De-energize crane using wall switch

Post-Operation Checklist (Powering down the crane)

The Lead Crane Operator must go through this checklist before considering work for the shift complete.  This checklist delineates the steps required to safely store the crane between uses.  

1. ( Load-block is in the fully up position (almost touching the upper limit chain)

2. ( Trolley has been moved as close to the wall / rotational axis of the beam as possible

3. ( Crane has been powered down using the OFF switch on the yellow hand-held controller box

4. ( Crane has been de-energized by turning off the switch at the wall, at the base of the crane

Training Procedures and Requirements
This section describes the training procedures to become an authorized Crane Operator at the 40m Lab.  The Authorized Crane Operator List should be updated after every training session.

1. Attend a crane safety training course offered by a certified instructor. 

2. Receive site-specific crane training at the 40m Lab

3. Accumulate a minimum of 20 hours of crane work under supervision of either Steve Vass or Bob Taylor

4. Authorized, experienced operators Steve Vass and Bob Taylor must have full confidence in trainee to become an operator before certification will be granted
